Аполипопротеин А1 ФС. Иммунотурбидиметрическое определение, биреагентная схема.

[image: image1.jpg]Item Name 5 |Apoat

Input range -9999.99 - 9999.99

DATA INFORMATIO!

CALIBRATION

TYPE _[spline E
STANDARD.

# o7 #
» 15 #
w32 #

NORMAL RANGF

MALE FEMALE

LOW | HIGH | LOW

HIGH

Serum

UNITS g1
DECIMALS |
ANALYSIS
TYPE eno E
Main W.Length 1 [s70 <
Sub Wilength 2 -
METHOD DiaSys Immuno
CORR.
SLOPE INTER
Y = X+ b

Urine.

Plasma

CSF

Dialysis

Other,

Page : 1

Next Page Save Return





Для калибровки рекомендуется использовать набор калибраторов TruCal Apo A1/B производства компании DiaSys.

Концентрации калибраторов указаны во вкладыше к набору калибраторов.

Контроль правильности измерений рекомендуется проводить по контрольной сыворотке TruLab L.

[image: image2.jpg]Item Name 5 ApoAl

SPAN B

FIRST
SECOND

T Clow

Select Single for single reagent, Double for double reagents system
ASPIRATION DATA PROCESS
KIND | Single © Double READ /ABSORBANCE LIMIT
START |_END
VOLUME o0 G o oW
SAMELH SUB_ 29 Bt W
REAGENTI VOL 750 il SHicHE]
REAGENT2 (50 |
EACTOR — | [enopoINTLIMIT |2
Blank correction|0.835 LINEAR CHECKEI
Th © oFF Con
Ri Blank |  Water.Blank (" R1.Blank1 Diluent | o9t © 10002
O PROZONE CHECK
START LIMIT(%)
0 LEVEL POINT [1

 High
 High

T “Lw

Page : 2

Print

Prev Page | [ NextPage Save Return





[image: image3.jpg]Item Name 5 ApoAl

Auto Rerun SW.

“ on « off

Auto Rerun Range (Result)

Con off  onf Off

Lower | Higher

Auto Rerun Condition (Absorbance)

Absorbance Range.

[ Lower | CoOnfof
[ Higher | C on® ofi
Prozone Range C on® off

Prev Page Save Return




