Аполипопротеин В ФС. Иммунотурбидиметрическое определение, биреагентная схема.

[image: image1.jpg]Item Name 6 |ApoB

Input range -9999.99 - 9999.99

DATA INFORMATIO!

CALIBRATION

TYPE [spiine B
STANDARD.

# o5 #
w |11 #
w |25 #

NORMAL RANGF

UNITS g1
DECIMALS |
ANALYSIS
TYPE eno E
Main W.Length 1 [s40 -
Sub Wilength2 100 =
METHOD DiaSys Immuno
CORR.
SLOPE INTER
Y= X+ @b

MALE FEMALE
LOW || HIGH | LOW | HIGH
Serum | [0.8 1.55 Jo.75 15
Urine
Plasma
cSF
Dialysis
Other

Next Page Save Return





Для калибровки рекомендуется использовать набор калибраторов TruCal Apo A1/B производства компании DiaSys.

Концентрации калибраторов указаны во вкладыше к набору калибраторов.

Контроль правильности измерений рекомендуется проводить по контрольной сыворотке TruLab L.

[image: image2.jpg]Item Name 6 ApoB

Select Single for single reagent, Double for. double reagents system

ASPIRATION DATA PROCESS
KIND | Single  Doubie READ ABSORBANCE LIMIT
START |_END
ARVOCUMER] WA 52 [ou ToWRG
SAMPLE |3 [ Low |
SUB_ 29 31 HGH |5
REAGENTA VOL 750 il ||
REAGENT2 (50 |
EACTOR ENDPOINT LIMIT [2
Blank correction|0.835 LINEAR CHECKEIT
Th © oFF Con
Ri Blank |  Water-Blank (" R1.Blank1 Diluent | o9t & 10002
I PROZONE CHECK
START LIMIT(%)
0 LEVEL POINT [1 FIRST
span |3 SECOND. [ [ Clow CHin

1 “low ( Hgh

Page : 2

Print

Prev Page | [ NextPage Save Return




[image: image3.jpg]Item Name 6 ApoB

Auto Rerun SW.

 on « off

Auto Rerun Range (Result)

C on off  Onf Off

Lower | | Higher

Auto Rerun Condition (Absorbance)

Absorbance Range.

[ Lower | CoOn®ofi
[ Higher | C on® off
Prozone Range C on® off

Serum

Urine

Plasma
csF

Dialysis

Other.

Prev Page Save Return




