oneHbA1c FS
Prestige 24i

Order information

Cat. No. 1 3329 .. .. …

Notes

1.
Please refer to the package insert for oneHbA1c FS for detailed information about the test on the following:
Clinical Relevance, Method and Principle, Composition and Stability of the Reagents, Specimens, Calibrators and Controls, Performance Characteristics regarding Measuring Range / Specificity/Interferences / Sensitivity/Limit of Detection / Precision (Reproducibility, Repeatability) / Method Comparison / Reference Ranges / Literature 

2. The stability of the reagent on board the analyser is at least one month provided that contamination and evaporation are avoided.

3. Manufactured by:  DiaSys Diagnostic Systems GmbH 
                              Alte Strasse 9, 65558 Holzheim, Germany.

NOTE: For optimal performance cuvette washes are recommended before and after this assay

	
	
	

	
	Item Name
	
	#
	
	HbA1c
	
	

	
	
	

	
	
	

	
	
	DATA INFORMATION
	
	CALIBRATION
	
	

	
	
	
	
	
	
	

	
	
	
	UNITS
	%
	
	
	
	TYPE
	Spline
	
	
	

	
	
	
	DECIMALS
	2
	
	
	
	STANDARD
	
	
	

	
	
	
	
	
	
	#1
	*
	#4
	*
	
	
	

	
	
	
	
	#2
	*
	#5
	
	
	
	

	
	
	ANALYSIS
	
	
	
	#3
	*
	#6
	
	
	
	

	
	
	
	
	
	
	
	

	
	
	
	TYPE
	END
	
	
	
	
	

	
	
	
	
	
	
	
	NORMAL RANGE
	
	

	
	
	
	Main W.Length 1
	660
	
	
	
	MALE
	FEMALE
	
	
	

	
	
	
	Main W.Length 2
	-
	
	
	
	LOW
	HIGH
	LOW
	HIGH
	
	
	

	
	
	
	
	
	
	
	
	Serum
	#
	#
	#
	#
	
	
	

	
	
	
	METHOD
	IT
	
	
	
	Urine
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	Plasma
	
	
	
	
	
	
	

	
	
	
	CSF
	
	
	
	
	
	
	

	
	
	
	Dialysis
	
	
	
	
	
	
	

	
	
	CORR.
	
	
	Other
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	SLOPE
	
	INTER
	
	
	
	
	
	
	
	
	
	

	
	
	
	Y =
	1
	X +
	0
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	

	
	
	

	
	Page : 1
	
	

	
	
	
	


	*input calibrator value
	

	# entered by the user

	
	


	
	
	

	
	Item Name
	
	#
	
	HbA1c
	
	

	
	
	
	
	
	
	
	

	
	
	

	
	
	ASPIRATION
	
	DATA PROCESS
	
	

	
	
	
	
	
	
	

	
	
	
	KIND
	( Single
	( Double
	
	
	READ
	
	ABSORBANCE LIMIT
	
	
	

	
	
	
	VOLUME
	
	
	
	
	START
	END
	
	
	
	
	
	

	
	
	
	SAMPLE
	5
	
	
	
	
	
	MAIN
	52
	54
	
	
	
	LOW
	0.000
	
	
	
	

	
	
	
	REAGENT1 VOL
	200
	µl
	
	
	
	
	SUB
	
	
	
	
	
	HIGH
	3.000
	
	
	
	

	
	
	
	REAGENT2
	100
	
	
	
	
	
	
	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	
	
	
	
	FACTOR
	
	ENDPOINT LIMIT
	2.000
	
	
	

	
	
	
	Third Mix.
	( OFF
	( ON
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	Blk correction
	1
	
	LINEAR CHECK(%)
	0
	
	
	

	
	
	
	R1 Blank
	( Water Blk
	( R1 Blank1
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	Dilution
	
	
	

	
	
	MONITOR
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	Diluent
	( 99:Dil1
	( 100:Dil2
	
	
	

	
	
	
	0 LEVEL POINT
	1
	
	
	
	
	
	
	
	
	
	

	
	
	
	SPAN
	3
	
	
	
	
	
	
	

	
	
	
	
	
	PROZONE CHECK
	
	
	

	
	
	
	
	
	START
	END
	LIMIT(%)
	
	
	
	

	
	
	
	
	FIRST
	
	
	
	
	
	
	

	
	
	
	
	SECOND
	
	
	
	( Low
	( High
	
	
	

	
	
	
	
	THIRD
	
	
	
	( Low
	( High
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	Page : 2
	
	
	

	
	
	
	
	


	
	
	

	
	Item Name
	
	#
	
	HbA1c
	
	

	
	
	

	
	
	

	
	
	Auto Rerun SW
	
	Auto Rerun Condition (Absorbance
	
	

	
	
	
	
	
	
	

	
	
	
	( On
	( Off
	
	
	
	Absorbance Range
	
	
	

	
	
	
	
	
	
	
	
	
	
	

	
	
	
	
	Lower
	( On
	( Off
	
	
	

	
	
	
	
	Higher
	( On
	( Off
	
	
	

	
	
	Auto Rerun Range (Result)
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	
	

	
	
	
	( On
	( Off
	( On
	( Off
	
	
	
	Prozone Range
	( On
	( Off
	
	
	

	
	
	
	Lower
	Higher
	
	
	
	
	

	
	
	
	Serum
	
	15.00
	
	
	
	

	
	
	
	Urine
	
	
	
	
	
	

	
	
	
	Plasma
	
	
	
	
	
	

	
	
	
	CSF
	
	
	
	
	
	

	
	
	
	Dialysis
	
	
	
	
	
	

	
	
	
	Other
	
	
	
	
	
	

	
	
	
	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	

	
	
	

	
	Page : 3
	
	

	
	
	
	


	*input calibrator value
	

	# entered by the user

	
	


845 3329 10 02 81

May 2007/1

