Холестерин ФС, ферментативный, монореагентный тест по конечной точке.

[image: image1.jpg]Item Name 16
Input range -9999.99 - 9999.99
DATA INFORMATION CALIBRATION
TYPE  [uincar =
UNITS mmolji
DECIMALS |2 STANDARD
#1 (109 “
ANALYSIS " #
" L
TYPE e =
Main W.Length 1 [s0s - HORMAL RN T
Sub Wilength2  [so0 - LOW | HIGH LOW | HIGH
Serum | 33 5.2 B3 [52
METHOD DiaSys PAP e | — ‘
Plasma [
CORR.. CSF I I
= SLOPE Pees ’n'"l Dialysis [
- [ Other ]
Page : 1 [[NextPage |[ save | [ Retun





Концентрация стандарта указана для мультикалибратора TruCal U производства компании DiaSys, лот № 5269

[image: image2.jpg]Item Name 16 Cho-D,

Select Single for single reagent, Double for double reagents system

ASPIRATION
KIND_ | & Singie. * Double

VOLUME
SAMPLE |3

REAGENT1 VOL 200 u
Third Mix. | = OFF oN
RiBlank | C Water-Blank © R1-Blank1

MONITOR

0 LEVEL POINT [1

DATA PROCESS
READ ABSORBANCE LIMIT
START || END
MAIN 52 54 Low fo
SUBE[ HIGH [3
EACTOR ENDPOINT LIMIT 2
Blank correction 1 LINEAR GHECKER

ut

n

Diluent | = 99:Dil1 100i12
PROZONE CHECK
START | END, LIMIT(S)

FIRST [

SPAN 5 SECOND ] g * Low High
THIRD [ [ Flow High
Prev Page | [ NextPage |[  Save | [ Retumn

Page : 2 Print





[image: image3.jpg]Item Name 16 Cho-D,

Input range -9999.99 - 9999.99

Auto Rerun SW. Auto Rerun Condition (Absorhance)
© Si0n Absorbance Range
Lower on © ofr
Auto Rerun Range (Resulq .
Auto Rerun Range (Resull —_— o
Prozone Range on © ofr
Serum f
Urine | [ [
Plasma | [ [
CSF [ [
Dialysis | | [ [
other | [ [
Page : 3 PrevPage | [ Save | [ Retun





