Ферритин ФС. Иммунотурбидиметрическое определение, биреагент.

[image: image1.jpg]| | §

Item Name 35 [FerD

Input range -9999.99 - 9999.99

DATA INFORMATION CALIBRATION
TYRE  [Logitz 5
UNITS mkg/l
DECIMALS [0 STANDARD.
# o | [0
ANALYSIS w200 #
w500 "
TYPE RaTE <

GF
Main W.length 1 [s00 NORMAL RANG!

Sub W.length 2

MALE FEMALE
LOW | HIGH  LOW | HIGH

Serum | [30 f300 10 [160

METHOD DiaSys Immuno i I
Plasma | [
s INTER co7 I S ——
SLUPE Dialysis | [
=5 X+ ® —
£ i b Other | [ I

Page : 1 [[NextPage |[ save | [ Retun




Для калибровки рекомендуется использовать набор калибраторов TruCal Ferritin производства компании DiaSys.

Концентрации калибраторов указаны во вкладыше к набору калибраторов.

Контроль правильности измерений рекомендуется проводить по контрольным сывороткам TruLab Protein, уровень 1 и уровень 2.

При переходе от второй страницы будет появляться сообщение о недостаточном объёме реагента. Из соображений экономии рекомендуется нажимать кнопку OK. При готовности к увеличению расхода реагентов на 20% во избежание появления подобных сообщений, на второй странице рекомендуется указывать следующие объёмы:

SAMPLE  10                 REAGENT1 VOL  200                REAGENT2  100

[image: image2.jpg]Item Name 3 FerrD,

Select Single for single reagent, Double for double reagents system

ASPIRATION
KIND. Single  Double

VOLUME
SAMPLE (8
REAGENT1VOL [160 ul
REAGENTZ g0

Third Mix. | © OFF oN
RiBlank | © Water-Blank © R1-Blank1
MONITOR

0 LEVEL POINT [1
seaN B

DATA PROCESS
READ ABSORBANCE LIMIT
START || END
MAINE5 Low o
DB HIGH 3
EACTOR ENDPOINT LIMIT [0
Blank correction [1 LINEAR CHECKEHS0
Di
Diluent | 99: 100012
PROZONE CHECK
START | END LIMITC)

FIRST
SECOND T | T F Low High
THIRD. | I T “low High

Page : 2 Print

Return

Prev Page | [ NextPage |[ Save





[image: image3.jpg]Item Name 3 FerrD,

Auto Rerun SW Auto Rerun Condition (Absorhance)
On * off Absorbance Range.
Lower on © ofr
Auto Rerun Range (Resulq -
Auto Rerun Range (Resull — o
on o | Con o
Lower Higher Prozone Range on © ofr

Serum | [ [

Urine |
Plasma | ([ [
CSF |
pialysis | [ [
Other. |

Page : 3 PrevPage | [ Save | [ Retun





